# ANNEX II + III:	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

# Contract title: Pprocurement of Mobile air Monitoring station with software and Trailer for project Green-Path

# p 1 /…

Publication reference: RORS00127-PP2 - TD08

Columns 1-2 should be completed by the contracting authority

Columns 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the contracting authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words ‘compliant’ or ‘yes’ are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

General Requirements

- All requirements stated and outlined in this document must be regarded as mandatory and the minimum acceptable criteria. All requirements outlined in this document are accompanied by the phrase "or equivalent".

- The tenderer is required to provide the specifications of the offered items in the Technical Offer, including details such as the manufacturer, product type, model. All documentation must be provided in English or Serbian, both in hardcopy and electronic formats.

- The selected contractor is required to provide all necessary documentation for the registration of the trailer in accordance with the relevant legal regulations in the Republic of Serbia.

| 1. Item number | 2. Specifications required | 3. Specifications offered | 4. Notes, remarks, ref to documentation | 5. Evaluation committee’s notes |
| --- | --- | --- | --- | --- |
| 1 | Mobile air Monitoring station, Quantity: 1 Unit |  |  |  |
|  | Required technical characteristics: Measurement performance – gases - Property NO2 Concentration range: 2.000 ppb Detection limit: 5 ppb - Property NO Concentration range: 2.000 ppb Detection limit: 5 ppb - Property O3 Concentration range: 2.000 ppb Detection limit: 5 ppb - Property CO Concentration range: 10.000 ppb Detection limit: 10 ppb - Property SO2 Concentration range: 0-20 ppb Detection limit: 10-100 ppb Field performance – gases - Property NO2 Correlation with reference: R: 0.90 Accuracy: 5 ppb Unit‐to‐unit correlation: R: 0.98 Precision: 3 ppb - Property NO Correlation with reference: R: 0.95 Accuracy: 8 ppb Unit‐to‐unit correlation: R: 0.96 Precision: 3 ppb - Property O3 Correlation with reference: R: 0.90 Accuracy: 6 ppb Unit‐to‐unit correlation: R: 0.95 Precision: 4 ppb - Property CO Correlation with reference: R: 0.85 Accuracy: 183 ppb Unit‐to‐unit correlation: R: 0.97 Precision: 25 ppb Measurement performance – environmental parameters - Humidity Accuracy for sensor element: 0–90 %RH: ±3 %RH 90–100 %RH: ±5 %RH Resolution: 0.1 %RH - Temperature Accuracy for sensor element: 0.3 °C (0.17 °F) at +20 °C (+68 °F) Resolution: 0.1 °C - Pressure (indicative) Accuracy: 15 hPa Resolution: 1 hPa Measurement performance - regulated particles - Property PM2.5 Concentration range: 0–1000 μg/m3 Detection limit: 0.1 μg/m3 - Property PM10 Concentration range: 0–2500 μg/m3 Detection limit: 0.1 μg/m3 Field performance – particles - Property PM2.5 Correlation with reference: R: 0.85 Accuracy: 3 μg/m3 Unit‐to‐unit correlation: R: 0.97 Precision: 2 μg/m3 - Property PM2.5 Correlation with reference: R: 0.85 Accuracy: 4 μg/m3 Unit‐to‐unit correlation: R: 0.97 Precision: 3 μg/m3 Measurement parameters • Air temperature, relative humidity, dew point, air pressure • NO2, NO, CO, CO2, O3, SO2 • PM2,5 ; PM10 • Global horizontal irradiance • Photovoltaic reference radiation, photovoltaic panel temperature, soiling ratio, transmission loss |  |  |  |
| 2 | Software air Monitoring station, Quantity: 1 Unit |  |  |  |
|  | Required technical characteristics: To collect data from all devices, to be able to define thresholds for parameters, email notifications, the ability to add devices, setup, status monitoring, configurable language, time zone, units, date/time format, availability for a combination of different measurement sources, and forecasting services. |  |  |  |
| 3 | Trailer for Mobile air Monitoring station, Quantity: 1 Unit |  |  |  |
|  | Required technical characteristics: (depending on the dimensions of the air quality measuring device, but in accordance with the dimensions of the device and the possibility of installation) Total mass (kg) 750 or more Useful space dimensions: minimum 205x132x39 Total dimensions minimum 312x178 Axle load capacity (kg) 750 Opening of the sides forward or backward or both Dismantling of all sides is optional Tarpaulin-lid is optional Spare wheel with carrier is optional The frame above the front side is optional Double-sided pages are optional Extension of ribbed wire, optional - height 34.5 cm |  |  |  |

| All items | Delivery of all items to following address: Provincial Secretariat for Urban Planning and Environmental Protection, Bulevar Mihajla Pupina, 16, 21101 Novi Sad, Republic of Serbia. |  |  |  |
| --- | --- | --- | --- | --- |
| All items | Testing of all basic functions of the instrument on a set of producers standard samples commonly used for the corresponding instrument. Installed equipment must be tested as system, compatible with existing system. |  |  |  |
| All items | Technical documentation for equipment (Operating manuals/ Users Guide/ Equipment operating instructions/ Cleaning procedures/ Maintenance procedures/ Calibration procedures) upon delivery. |  |  |  |
| All items | Warranty Warranty period 365 days from the issuance of PAC (Provisional Acceptance Certificate) in accordance with article 32 SC and GC of the Contract. Offer must include warranty service description including: - Service organisation contact data including name, postal address, telephone number, fax number and e-mail address; - Guaranteed that any requests for services will be attended to within 48 hours; - Guarantee that all items can be repaired or alternatively replaced within a maximum of 7 days; - Letter of confirmation that genuine spare parts and consumables will be available for a period of minimum 3 years from the date of final acceptance of the equipment. |  |  |  |
| All items | Commercial warranty 365 days from the issuance of FAC (Final Acceptance Certificate) in accordance with the conditions laid down in Article 32 of the General Conditions and Article 33 of the Special Conditions. Detailed description of the organisation of the proposed service and description of the Manufacturer’s commercial warranty shall be included in the offer. |  |  |  |

<Signature of authorised representative of the legal entity >

< Name and position of authorised representative of the legal entity >

< Date >